## Hoja1
NDIR CO2 sensors for indoor air quality monitoring system
Last update: 7/2/2021
Manufacturer | Model (article number) | Effective range | Accuracy | Interface | Power supply | Average current | For battery project | Preheat time | Response time | Minimum lifespan | Sensors | Price end user (excluded shipping) | Product page | Datasheet/Manual | Shop 1 (end user) | Shop 2 (end user) | Shop 3 (end user)
Amphenol Advanced Sensors | T6713 | 400-2000 ppm | ± (25 ppm + 3% reading value) | UART, I2C, PWM | 4.5-5.5 VDC | 20 mA |  | < 120 s | < 180 s (90%) | 15 years | CO2 | 48.19 - 70.99 € | https://www.amphenol-sensors.com/en/telaire/co2/525-co2-sensor-modules/3399-t6713 | https://www.amphenol-sensors.com/en/literature/276-telaire/4-co2/5-t6700/297-telaire-t6713-miniature-co2-sensor-module-datasheet | https://www.mouser.es/ProductDetail/Amphenol-Advanced-Sensors/T6713/?qs=N10xTRV9dj26U%252BpBD9TY%2FA%3D%3D | https://www.digikey.es/product-detail/en/amphenol-advanced-sensors/T6713/235-1373-ND/5027891 | https://es.aliexpress.com/item/32815695531.html
Amphenol Advanced Sensors | T6713-5K | 400-5000 ppm | ± (30 ppm + 3% reading value) | UART, I2C, PWM | 4.5-5.5 VDC | 20 mA |  | < 120 s | < 180 s (90%) | 15 years | CO2 | 51.46 - 70.99 € | https://www.amphenol-sensors.com/en/telaire/co2/525-co2-sensor-modules/3399-t6713 | https://www.amphenol-sensors.com/en/literature/276-telaire/4-co2/5-t6700/297-telaire-t6713-miniature-co2-sensor-module-datasheet | https://www.mouser.es/ProductDetail/Amphenol-Advanced-Sensors/T6713-5k/?qs=N10xTRV9dj2DLVnP5YzNbg%3D%3D | https://www.digikey.es/product-detail/en/amphenol-advanced-sensors/T6713-5K/235-1374-ND/5027892
Cubic | CM1106-C | 400-5000 ppm | ± (50 ppm + 5% reading value) | UART, I2C, PWM | 4.5-5.5 VDC | < 45 mA |  |  | ≤ 30 s | 10 years | CO2 |  | https://en.gassensor.com.cn/CO2Sensor/info_itemid_86.html | https://en.gassensor.com.cn/Product_files/Specifications/CM1106-C%20Single%20Beam%20NDIR%20CO2%20Sensor%20Module%20Specification.pdf
Cubic | CM1106SL-NS | 400-5000 ppm | ± (50 ppm + 5% reading value) | UART | 3.5-5.5 VDC | 30 uA (2 min.) | Yes |  | 120 s | 15 years | CO2 |  | https://en.gassensor.com.cn/CO2Sensor/info_itemid_930.html | https://en.gassensor.com.cn/Product_files/Specifications/CM1106SL-N%20Super%20Low%20Power%20CO2%20Sensor%20Module%20Specification.pdf
Gas Sensing Solutions | COZIR-AH-E-1 | 400-10000 ppm | ± (30 ppm + 3% reading value) | UART | 3.25-5.5 VDC | 15 mA |  |  | 30 s (90%) |  | CO2, Humidity, Temperature | 73.78 - 100.14 € | https://www.gassensing.co.uk/product/cozir-a-co2-sensor/ | https://www.gassensing.co.uk/wp-content/uploads/2021/02/CozIR-A-Data-Sheet-Rev-4.4.pdf | https://www.digikey.es/product-detail/en/gas-sensing-solutions-ltd/COZIR-AH-E-1/2091-COZIR-AH-E-1-ND/13148756 | https://es.rs-online.com/web/p/circuitos-integrados-de-sensores-ambientales/1809025/
Gas Sensing Solutions | COZIR-LP-5000 | 400-5000 ppm | ± (30 ppm + 3% reading value) | UART | 3.25-5.5 VDC | 15 mA |  |  | 30 s (90%) | 15 years | CO2 | 63.41 - 86.06 € | https://www.gassensing.co.uk/product/cozir-lp-co2-sensor/ | https://www.gassensing.co.uk/wp-content/uploads/2020/07/CozIR-LP-Data-Sheet-Rev-4.4.pdf | https://www.digikey.es/product-detail/es/gas-sensing-solutions-ltd/COZIR-LP-5000/2091-COZIR-LP-5000-ND/9952877 | https://es.rs-online.com/web/p/circuitos-integrados-de-sensores-ambientales/1809024/
Honeywell | CRIR M1 (C06-0801-000) | 400-2000 ppm | ± (40 ppm + 3% reading value) | UART | 4.5-5.5 VDC | 30 mA |  | 180 s | ≤ 120 s | 10 years | CO2 | 39.48 - 51.34 € | https://sps.honeywell.com/us/en/products/sensing-and-iot/sensors/gas-sensors/crir-series | https://prod-edam.honeywell.com/content/dam/honeywell-edam/sps/siot/en-us/products/sensors/gas-sensors/crir-series/documents/sps-siot-carbon-dioxide-crir-m1-sensor-datasheet-000839-ciid-179497.pdf | https://www.arrow.com/en/products/c06-0801-000/honeywell
SenseAir | S8 LP (004-0-0053) | 400-2000 ppm | ± (40 ppm + 3% reading value) | UART, PWM | 4.5-5.25 VDC | 18 mA |  |  | < 120 s (90%) | 15 years | CO2 | 23.17 - 36.42 € | https://senseair.com/products/size-counts/s8-lp/ | https://rmtplusstoragesenseair.blob.core.windows.net/docs/publicerat/PSP126.pdf | https://es.aliexpress.com/item/4000783293177.html | https://www.digikey.es/product-detail/es/senseair-north-america-inc/004-0-0053/2194-004-0-0053-ND/10416532
SenseAir | Sunrise (006-0-0002) | 400-5000 ppm | ± (30 ppm + 3% reading value) | UART, I2C | 3.05-5.5 VDC | 38 uA | Yes |  | 16 s | 15 years | CO2 | 44.84 € | https://senseair.com/products/power-counts/sunrise-hvac/ | https://rmtplusstoragesenseair.blob.core.windows.net/docs/Dev/publicerat/PSP4731.pdf | https://www.digikey.es/product-detail/es/senseair-north-america-inc/006-0-0002/2194-006-0-0002-ND/10416535 | https://www.alibaba.com/product-detail/006-0-0002-industrial-SENSOR-006_1600146627614.html
Sensirion | SCD-41 (SEK-SCD41) | 400-5000 ppm | ± (40 ppm + 5% reading value) | I2C | 2.4-5.5 VDC | 2.8-18 mA |  |  | 60 s (63%) | 10 years | CO2, Humidity, Temperature | 42.94 € | https://www.sensirion.com/en/environmental-sensors/carbon-dioxide-sensors/carbon-dioxide-sensor-scd4x/ | https://www.sensirion.com/fileadmin/user_upload/customers/sensirion/Dokumente/9.5_CO2/Sensirion_CO2_Sensors_SCD4x_Datasheet.pdf | https://www.mouser.es/ProductDetail/Sensirion/SEK-SCD41-Sensor/?qs=HBWAp0VN4RiEPPmz5KP5eQ%3D%3D
Sensirion | SCD-30 (1-101625-10) | 400-10000 ppm | ± (30 ppm + 3% reading value) | UART, I2C, PWM | 3.3-5.5 VDC | 19 mA |  |  | 20 s (63%) | 15 years | CO2, Humidity, Temperature | 30.57 - 44.47 € | https://www.sensirion.com/en/environmental-sensors/carbon-dioxide-sensors/carbon-dioxide-sensors-co2/ | https://www.sensirion.com/fileadmin/user_upload/customers/sensirion/Dokumente/9.5_CO2/Sensirion_CO2_Sensors_SCD30_Datasheet.pdf | https://es.aliexpress.com/item/1005001392172293.html | https://www.mouser.es/ProductDetail/Sensirion/SCD30?qs=rrS6PyfT74fdywu4FxpYjQ== | https://www.digikey.es/products/es/sensors-transducers/gas-sensors/530?k=scd-30
Winsen | MH-Z19B | 400-5000 ppm | ± (50 ppm + 5% reading value) | UART, PWM | 4.5-5.5 VDC | < 20 mA |  | 180 s | < 120 s (90%) | 5 years | CO2 | 18.61 € | https://www.winsen-sensor.com/sensors/co2-sensor/mh-z19b.html | https://www.winsen-sensor.com/d/files/infrared-gas-sensor/ndir-co2-sensor/mh-z19b-co2-manual(ver1_6).pdf | https://es.aliexpress.com/item/1005001865093513.html
Winsen | MH-Z19C | 400-5000 ppm | ± (50 ppm + 5% reading value) | UART, PWM | 4.9-5.1 VDC | < 40 mA |  | 60 s | < 120 s (90%) | 5 years | CO2 | 17.09 € | https://www.winsen-sensor.com/sensors/co2-sensor/mh-z19c.html | https://www.winsen-sensor.com/d/files/infrared-gas-sensor/mh-z19c-pins-type-co2-manual-ver1_0.pdf | https://es.aliexpress.com/item/1005001947070873.html
ZyAura (Radiant) | ZG09 | 400-10000 ppm | ± (50 ppm + 3% reading value) | UART, I2C, PWM | 4.5-5.5 VDC | < 34 mA |  | 30 s | 60s |  | CO2 | 27.11 € | http://www.zyaura.com/products/ZG09.asp | http://www.zyaura.com/support/manual/pdf/ZyAura_CO2_Monitor_Carbon_Dioxide_ZG09%20English%20user%20manual_1808.pdf | https://es.aliexpress.com/item/4000215863926.html
Discontinued products:
Amphenol Advanced Sensors | T6703-5K | 400-5000 ppm | ± 75 ppm or 10% reading (whichever is greater) | UART, I2C, PWM | 4.5-5.5 VDC | 20 mA |  | < 120 s | < 180 s (90%) | 10 years | CO2 | 21.41 € | https://www.amphenol-sensors.com/en/telaire/co2/525-co2-sensor-modules/3400-t6703 | https://www.amphenol-sensors.com/en/services-and-support/25-telaire-t6703-series-miniature-co2-sensor-modules-datasheet | https://es.aliexpress.com/item/4001009559168.html
Cubic | CM1106 | 400-2000 ppm | ± (50 ppm + 5% reading value) | UART, PWM | 4.75-5.25 VDC | 70 mA |  |  | < 120 s (90%) | 8 years | CO2 | 3.27 € | N/A | http://80.76.226.157/data/files/carbon_dioxide/CM1106%20CO2%20SENSOR%20MODULE%20INTRODUCTION.pdf | https://es.aliexpress.com/item/4001063549786.html
Figaro | CDM7160-C00 | 300-5000 ppm | ± (50 ppm + 3% reading value) | UART, I2C, PWM | 4.75-5.25 VDC | 10 mA |  |  | 120 s (90%) |  | CO2 | 44.55-51.73 € | https://www.figaro.co.jp/en/product/docs/cdm7160-c00_product%20infomation(en)_rev00.pdf (Discontinuance reason: https://www.figaro.co.jp/en/topic/2020/11/cdm7160-2.html)
Notes:
* Below 400 ppm measure was excluded in range, some sensors provide measure but it isn't effective.
* Accuracy indicated is the typical but the error can be a little more, especially for values more than 2000 ppm.
* Precision for all these sensors is 1 ppm but the interesting is the accuracy
* For power supply you have to consider peak consumption, see datasheet of the manufacturer